## Slide 1
Validation of a task-specific deep-learning-based denoising approach for myocardial perfusion SPECT Md Ashequr Rahman2, Zezhang Yang1, Zitong Yu2, Barry A. Siegel3 and Abhinav K. Jha1,2,3 1Department of Electronic and System Engineering, Washington University in St. Louis, St. Louis, MO, USA
2Department of Biomedical Engineering, Washington University in St. Louis, St. Louis, MO, USA
3Mallinckrodt Institute of Radiology, Washington University in St. Louis, St. Louis, MO, USA
1. Introduction
Important need for methods to process myocardial perfusion imaging (MPI) SPECT images at lower dose/acquisition time1.
Deep-learning-based methods have been developed for this goal, but they have shown limited efficacy on clinical task of defect detection3.
We recently developed DEMIST: A task-specific deep-learning-based denoising method to process low-count MPI-SPECT scans4. Here, we expand on the validation of that technique.
Human visual system process images through frequency selective channels5. Motivated by this fact, channelized Hotelling observers (CHOs) have been proposed.
2. Task-specific deep-learning-based denoising method (DEMIST)
Frequency selective channels
Channel vector,
Apply Hotelling template,
Scalar test statistic,
Image
Financial support: NIBIB R01-EB031051, R01-EB031962, NSF CAREER award || Contact information: y.zezhang@wustl.edu || Lab website: jhalab.wustl.edu
Proposed denoising method
Background: Channel mechanism of observer
CHO with rotationally symmetric square channels was shown to emulate human-observer performance6.
3. Evaluation of the DEMIST methodology
4. Conclusion
As per RELAINCE guidelines7, the following claim is derived for DEMIST: DEMIST, a deep-learning-based denoising method for myocardial perfusion SPECT yielded statistically significant improved performance compared to the low-dose images on the task of detecting perfusion defects as evaluated using a retrospective clinical study with single-center data and with an anthropomorphic model observer.
Further, we observe in qualitative evaluations that DEMIST also has improved sensitivity and specificity compared to low-dose images The results motivate further evaluation of the method.
5. References
Study design: Data collection and curation
6.25% and 12.5% denote the low-dose levels.
Normal-dose image
Low-dose image
Proposed task-specific DL denoising
CHO-based model-observer evaluation on defect-detection task
Evaluation framework
ROC analysis to determine AUC
Denoised image
Evaluation results
The DEMIST method significantly improved performance on defect-detection task compared to low-dose images for defects located in anterior and inferior walls of left ventricle (a) over entire population and (b) for all severities considered at 6.25% dose level and (c) for all defect extent.
Qualitative evaluation: Sensitivity of DEMIST
DEMIST denoising method tend to wash out the defect. The DEMIST method was able to preserve the defect contrast.
Normal dose
Low dose
Task-agnostic DL denoising
DEMIST
Qualitative evaluation: Specificity of DEMIST
Normal dose
Low dose
DEMIST
We observe that multiple false defects are visible in the low-dose scans. DEMIST either removes or mitigates these false defects, resulting in improved specificity compared to low-dose images.
Loss function,
Minimizes the mean square error between true normal-dose and predicted normal-dose image
Minimizes the mean square error of channel vectors between true normal-dose and predicted normal-dose image
ROC analysis was performed on the test statistics derived with the CHO. The area under the empirical ROC curve (AUC) was used as the figure of merit.
Beller et al, J. Nuc. Card. 2010
“Heart SPECT Imaging” by Patrick J. Lynch, licensed under CC BY 2.5, adapted
Li et al., Med. Phys. 2023
Rahman et al., ArXiv. 2023
5. Barrett & Myers, Found. of Image Science 2013
6. Sankaran et al., J. Nucl. Med. 2002
7. Jha et al., J. Nucl. Med. 2022
(a)
(b)
(c)